## Metadata
Property | Value
URL | http://hl7.org/fhir/StructureDefinition/DeviceAlert
Version | 6.0.0-ballot4
Name | DeviceAlert
Title
Status | active
Experimental | false
Date | 2026-06-17T19:53:47+00:00
Publisher | HL7 International / Devices
Contact | null (http://www.hl7.org/Special/committees/fiwg)
Contact | null (http://www.hl7.org/Special/committees/healthcaredevices)
Jurisdiction | World
Description | Describes a physiological or technical alert condition report originated by a device. The DeviceAlert resource is derived from the ISO/IEEE 11073-10201 Domain Information Model standard, but is more widely applicable.
Purpose
Copyright
FHIR Version | 6.0.0-ballot4
Kind | resource
Type | DeviceAlert
Base Definition | http://hl7.org/fhir/StructureDefinition/DomainResource
Abstract | false
Derivation | specialization

## Elements
ID | Path | Slice Name | Alias(s) | Label | Min | Max | Must Support? | Is Modifier? | Is Summary? | Type(s) | Short | Definition | Comments | Requirements | Default Value | Meaning When Missing | Fixed Value | Pattern | Example | Minimum Value | Maximum Value | Maximum Length | Binding Strength | Binding Description | Binding Value Set | Code | Slicing Discriminator | Slicing Description | Slicing Ordered | Slicing Rules | Base Path | Base Min | Base Max | Condition(s) | Constraint(s) | Mapping: Workflow Pattern | Mapping: FiveWs Pattern Mapping | Mapping: ISO/IEEE 11073-10207: Domain Information and Service Model for Service-Oriented Point-of-Care Medical Device Communication | Mapping: ISO/IEEE 11073-10201: Domain Information Model | Mapping: RIM Mapping
DeviceAlert | DeviceAlert |  |  |  | 0 | * |  |  |  |  | Documentation of an alert (a.k.a. alarm) generated by a device indicating a noteworthy condition | Describes a noteworthy condition or occurrence determined to exist by a device. | The DeviceAlert resource is derived from the ISO/IEEE 11073-10201 Domain Information Model standard, but is more widely applicable. | [17 empty] | DeviceAlert | 0 | * |  | dom-2:If the resource is contained in another resource, it SHALL NOT contain nested Resources {contained.contained.empty()} dom-3:If the resource is contained in another resource, it SHALL be referred to from elsewhere in the resource or SHALL refer to the containing resource {contained.where((('#'+id.trace('id') in %resource.descendants().select(reference ¦ as(uri))) or descendants().where(reference='#' ¦ as(uri)='#').exists()).not()).trace('unmatched', id).empty()}dom-4:If a resource is contained in another resource, it SHALL NOT have a meta.versionId or a meta.lastUpdated {contained.meta.versionId.empty() and contained.meta.lastUpdated.empty()}dom-5:If a resource is contained in another resource, it SHALL NOT have a security label {contained.meta.security.empty()}dom-6:A resource should have narrative for robust management {text.`div`.exists()} | Event | administrative.device | AlertConditionDescriptor, AlertSignalState, AlertConditionState | AlertCondition, AlertState | Entity, Role, or Act
DeviceAlert.id | DeviceAlert.id |  |  |  | 0 | 1 |  |  | Y | string | Logical id of this artifact | The logical id of the resource, as used in the URL for the resource. Once assigned, this value never changes. | Within the context of the FHIR RESTful interactions, the resource has an id except for cases like the create and conditional update. Otherwise, the use of the resouce id depends on the given use case. | [17 empty] | Resource.id | 0 | 1
DeviceAlert.meta | DeviceAlert.meta |  |  |  | 0 | 1 |  |  | Y | Meta | Metadata about the resource | The metadata about the resource. This is content that is maintained by the infrastructure. Changes to the content might not always be associated with version changes to the resource. | [18 empty] | Resource.meta | 0 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())}
DeviceAlert.implicitRules | DeviceAlert.implicitRules |  |  |  | 0 | 1 |  | Y | Y | uri | A set of rules under which this content was created | A reference to a set of rules that were followed when the resource was constructed, and which must be understood when processing the content. Often, this is a reference to an implementation guide that defines the special rules along with other profiles etc. For DomainResources, if more than one implicitRules policy applies, additional policies may be conveyed using the [additionalImplicitRules](https://build.fhir.org/ig/HL7/fhir-extensions/StructureDefinition-additionalImplicitRules.html) extension. | Asserting this rule set restricts the content to be only understood by a limited set of trading partners. This inherently limits the usefulness of the data in the long term. However, the existing health eco-system is highly fractured, and not yet ready to define, collect, and exchange data in a generally computable sense. Wherever possible, implementers and/or specification writers should avoid using this element. Often, when used, the URL is a reference to an implementation guide that defines these special rules as part of its narrative along with other profiles, value sets, etc. | [17 empty] | Resource.implicitRules | 0 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())}
DeviceAlert.language | DeviceAlert.language |  |  |  | 0 | 1 |  |  |  | code | Language of the resource content | The base language in which the resource is written. | Language is provided to support indexing and accessibility (typically, services such as text to speech use the language tag). The html language tag in the narrative applies to the narrative. The language tag on the resource may be used to specify the language of other presentations generated from the data in the resource. Not all the content has to be in the base language. The Resource.language should not be assumed to apply to the narrative automatically. If a language is specified, it should it also be specified on the div element in the html (see rules in HTML5 for information about the relationship between xml:lang and the html lang attribute). | [9 empty] | required | IETF language tag for a human language | http://hl7.org/fhir/ValueSet/all-languages¦6.0.0-ballot4 | [5 empty] | Resource.language | 0 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())}
DeviceAlert.text | DeviceAlert.text |  | narrative htmlxhtmldisplay |  | 0 | 1 |  |  |  | Narrative | Text summary of the resource, for human interpretation | A human-readable narrative that contains a summary of the resource and can be used to represent the content of the resource to a human. The narrative need not encode all the structured data, but is required to contain sufficient detail to make it "clinically safe" for a human to just read the narrative. Resource definitions may define what content should be represented in the narrative to ensure clinical safety. | Contained resources do not have a narrative. Resources that are not contained SHOULD have a narrative. In some cases, a resource may only have text with little or no additional discrete data (as long as all minOccurs=1 elements are satisfied). This may be necessary for data from legacy systems where information is captured as a "text blob" or where text is additionally entered raw or narrated and encoded information is added later. | [17 empty] | DomainResource.text | 0 | 1 | dom-6 | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  |  |  |  | Act.text?
DeviceAlert.contained | DeviceAlert.contained |  | inline resources anonymous resourcescontained resources |  | 0 | * |  |  |  | Resource | Contained, inline Resources | These resources do not have an independent existence apart from the resource that contains them - they cannot be identified independently, nor can they have their own independent transaction scope. This is allowed to be a Parameters resource if and only if it is referenced by a resource that provides context/meaning. | This should never be done when the content can be identified properly, as once identification is lost, it is extremely difficult (and context dependent) to restore it again. Contained resources may have profiles and tags in their meta elements, but SHALL NOT have security labels. Contained resources may be a resource type defined in the FHIR specification, or an [additional resource](resource.html#additional). | [17 empty] | DomainResource.contained | 0 | * | dom-2 dom-4dom-3dom-5 | [5 empty] | N/A
DeviceAlert.extension | DeviceAlert.extension |  | extensions user content |  | 0 | * |  |  |  | Extension | Additional content defined by implementations | May be used to represent additional information that is not part of the basic definition of the resource. To make the use of extensions safe and managable, there is a strict set of governance applied to the definition and use of extensions. Though any implementer can define an extension, there is a set of requirements that SHALL be met as part of the definition of the extension. | There can be no stigma associated with the use of extensions by any application, project, or standard - regardless of the institution or jurisdiction that uses or defines the extensions. The use of extensions is what allows the FHIR specification to retain a core level of simplicity for everyone. | [17 empty] | DomainResource.extension | 0 | * |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  |  |  |  | N/A
DeviceAlert.modifierExtension | DeviceAlert.modifierExtension |  | extensions user content |  | 0 | * |  | Y | Y | Extension | Extensions that cannot be ignored | May be used to represent additional information that is not part of the basic definition of the resource and that modifies the understanding of the element that contains it and/or the understanding of the containing element's descendants. Usually modifier elements provide negation or qualification. To make the use of extensions safe and managable, there is a strict set of governance applied to the definition and use of extensions. Though any implementer is allowed to define an extension, there is a set of requirements that SHALL be met as part of the definition of the extension. Applications processing a resource are required to check for modifier extensions. Modifier extensions SHALL NOT change the meaning of any elements on Resource or DomainResource (including cannot change the meaning of modifierExtension itself). | There can be no stigma associated with the use of extensions by any application, project, or standard - regardless of the institution or jurisdiction that uses or defines the extensions. The use of extensions is what allows the FHIR specification to retain a core level of simplicity for everyone. | Modifier extensions allow for extensions that *cannot* be safely ignored to be clearly distinguished from the vast majority of extensions which can be safely ignored. This promotes interoperability by eliminating the need for implementers to prohibit the presence of extensions. For further information, see the [definition of modifier extensions](extensibility.html#modifierExtension). | [16 empty] | DomainResource.modifierExtension | 0 | * |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  |  |  |  | N/A
DeviceAlert.identifier | DeviceAlert.identifier |  |  |  | 0 | * |  |  | Y | Identifier | Business identifier for this device alert | Business identifiers assigned to this alert, by the source device, gateway software, manufacturers, or other systems or organizations. These identifiers remain constant as the resource is updated and propagates from server to server. | Note: This is a business identifier, not a resource identifier (see [discussion](resource.html#identifiers)). It is best practice for the identifier to only appear on a single resource instance, however business practices may occasionally dictate that multiple resource instances with the same identifier can exist - possibly even with different resource types. For identifiers assigned to an alert by the device or gateway software, the `system` element of the identifier might be set to an identifier of the device. | Allows identification of the device alert as it is known by various participating systems and in a way that remains consistent across servers. | [16 empty] | DeviceAlert.identifier | 0 | * |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} | Event.identifier | FiveWs.identifier | AlertOccurrenceIdentifier | AlertCondition/alert-info-id
DeviceAlert.procedure | DeviceAlert.procedure |  |  |  | 0 | * |  |  | Y | Reference(Procedure) | Procedure during which the alert occurred | The procedure (or procedures) during which the alert condition was raised. | [18 empty] | DeviceAlert.procedure | 0 | * |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} | Event.partOf
DeviceAlert.status | DeviceAlert.status |  |  |  | 1 | 1 |  | Y | Y | code | in-progress ¦ completed ¦ entered-in-error ¦ unknown | The current state of the device alert. | Unknown does not represent "other" - one of the defined statuses must apply. Unknown is used when the authoring system is not sure what the current status is. | [9 empty] | required | in-progress ¦ completed ¦ entered-in-error ¦ unknown | http://hl7.org/fhir/ValueSet/devicealert-status¦6.0.0-ballot4 | [5 empty] | DeviceAlert.status | 1 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} | Event.status | FiveWs.status |  | AlertState
DeviceAlert.category | DeviceAlert.category |  |  |  | 0 | * |  |  | Y | CodeableConcept | High level categorization of device alert | Partitions the device alert into one or more categories that can be used to filter searching, to govern access control and/or to guide system behavior. | Categorization might be done automatically (inferred by code) or manually by user assertion. The absence of a category may limit the ability to determine when the element should be handled, so strong consideration should be given to how systems will be able to determine category values for legacy data and how data that cannot be categorized will be handled. As well, some categories might not be mutually exclusive, so systems should prepare for multiple declared categories - even within a single category 'axis'. | [9 empty] | example |  | http://terminology.hl7.org/ValueSet/devicealert-category | [5 empty] | DeviceAlert.category | 0 | * |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} | Event.category | FiveWs.class
DeviceAlert.type | DeviceAlert.type |  |  |  | 0 | 1 |  |  |  | CodeableConcept | The kind of alert | The type of alert, such as whether it is related to the patient condition; to irregular device state; or to workflow or regular device state, such as timers. | [10 empty] | extensible |  | http://terminology.hl7.org/ValueSet/devicealert-type | [5 empty] | DeviceAlert.type | 0 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  | FiveWs.class | AlertConditionDescriptor/Kind | AlertCondition/alert-type
DeviceAlert.priority | DeviceAlert.priority |  |  |  | 0 | 1 |  |  | Y | CodeableConcept | high ¦ medium ¦ low ¦ info | The alert priority is usually reported by the source device. A priority of `info` SHOULD indicate that the alert is “for information only” and urgent action is not required. The element MAY be omitted if the priority is unknown. | [10 empty] | extensible |  | http://hl7.org/fhir/ValueSet/devicealert-priority | [5 empty] | DeviceAlert.priority | 0 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  | FiveWs.grade | If present, AlertConditionState/ActualPriority; otherwise AlertConditionDescriptor/Priority | AlertCondition/alert-type
DeviceAlert.code | DeviceAlert.code |  |  |  | 1 | 1 |  |  | Y | CodeableConcept | The meaning of the alert | A code that indicates the specific condition that triggered the alert. | [10 empty] | preferred |  | http://terminology.hl7.org/ValueSet/devicealert-condition | [5 empty] | DeviceAlert.code | 1 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} | Event.code | FiveWs.what[x] | AlertConditionDescriptor/Type | AlertCondition/alert-code
DeviceAlert.subject | DeviceAlert.subject |  |  |  | 1 | 1 |  |  | Y | Reference(Patient¦Device¦BiologicallyDerivedProduct¦Group¦Location¦Medication¦NutritionProduct¦Specimen) | Who or what the alert is about | Who or what the alert is about. |  | Links the device alert to the appropriate subject context. May also affect access control. | [16 empty] | DeviceAlert.subject | 1 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} | Event.subject | FiveWs.subject
DeviceAlert.encounter | DeviceAlert.encounter |  |  |  | 0 | 1 |  |  | Y | Reference(Encounter) | Encounter during which the alert condition occurred | The Encounter during which the alert condition was raised. |  | Links the device alert to the Encounter context. May also affect access control. | [16 empty] | DeviceAlert.encounter | 0 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} | Event.encounter | FiveWs.context
DeviceAlert.presence | DeviceAlert.presence |  |  |  | 1 | 1 |  |  | Y | boolean | Whether the alert condition is currently active | Whether the alert condition is currently active. | [18 empty] | DeviceAlert.presence | 1 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  |  | AlertConditionState/Presence
DeviceAlert.occurrence[x] | DeviceAlert.occurrence[x] |  |  |  | 0 | 1 |  |  | Y | dateTime Period | When the alert condition occurred/is occurring | This element is used to record the date or time period when the alert condition did occur or is occurring. | This element is used to record the time period during which the alert condition is or was active. The status code allows differentiation of whether the timing reflects a historic event or an ongoing event. A date/time value SHOULD be used for instantaneous conditions. Partial date/time values are ambiguous about whether they represent the condition occurring at some point in the partially specified period, or for all of the period, and (absent additional guidance) SHOULD be avoided. Ongoing events SHOULD not include a Period upper bound. | [17 empty] | DeviceAlert.occurrence[x] | 0 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} | Event.occurrence[x] | FiveWs.done[x] | AlertConditionState/DeterminationTime
DeviceAlert.device | DeviceAlert.device |  |  |  | 0 | 1 |  |  | Y | Reference(Device¦DeviceMetric) | The Device (or DeviceMetric) that detected the alert condition | Indicates the device that detected the alert condition. The device could be a top-level Device or component Device (such as an MDS, VMD, or Channel); or might identify the specific DeviceMetric of a Device (e.g., a heart rate reading) that was in an alert condition. | [18 empty] | DeviceAlert.device | 0 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} | Event.product | FiveWs.author | AlertConditionDescriptor/Source | AlertCondition/alert-source
DeviceAlert.acknowledged | DeviceAlert.acknowledged |  |  |  | 0 | 1 |  |  |  | boolean | Whether the alert condition has been acknowledged | An indication of whether (true) or not (false) the alert condition has been acknowledged. |  |  |  | If absent, no inference SHOULD be made about whether the alert condition has been acknowledged. | [14 empty] | DeviceAlert.acknowledged | 0 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())}
DeviceAlert.acknowledgedBy | DeviceAlert.acknowledgedBy |  |  |  | 0 | 1 |  |  |  | Reference(Patient¦Device¦Practitioner¦PractitionerRole¦RelatedPerson) | Who acknowledged the alert condition | An indication of who or what acknowledged the alert condition. | [18 empty] | DeviceAlert.acknowledgedBy | 0 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} | Event.actor | FiveWs.actor
DeviceAlert.location | DeviceAlert.location |  |  |  | 0 | 1 |  |  | Y | Reference(Location) | Location of the subject when the alert was raised | The principal physical location of the subject at the time the alert condition occurred. This could be different from the location of the alerting device at that time, and from the current location of either the subject or the alert condition detecting device. |  | Provides context around the alert occurrence (e.g. if it occurred inside or outside a dedicated healthcare setting). | [16 empty] | DeviceAlert.location | 0 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} | Event.location | FiveWs.where[x]
DeviceAlert.derivedFrom | DeviceAlert.derivedFrom |  |  |  | 0 | * |  |  |  | BackboneElement | The value causing the alert condition | The value causing the alert condition. | [18 empty] | DeviceAlert.derivedFrom | 0 | * |  | ele-1:All FHIR elements must have a @value or children or both {hasValue() or (children().count() > id.count())} |  | FiveWs.why[x]
DeviceAlert.derivedFrom.id | DeviceAlert.derivedFrom.id |  |  |  | 0 | 1 |  |  |  | string | Unique id for inter-element referencing | Unique id for the element within a resource (for internal references). This may be any string value that does not contain spaces. | [18 empty] | Element.id | 0 | 1 | ele-1 | [5 empty] | n/a
DeviceAlert.derivedFrom.extension | DeviceAlert.derivedFrom.extension |  | extensions user content |  | 0 | * |  |  |  | Extension | Additional content defined by implementations | May be used to represent additional information that is not part of the basic definition of the element. To make the use of extensions safe and managable, there is a strict set of governance applied to the definition and use of extensions. Though any implementer can define an extension, there is a set of requirements that SHALL be met as part of the definition of the extension. | There can be no stigma associated with the use of extensions by any application, project, or standard - regardless of the institution or jurisdiction that uses or defines the extensions. The use of extensions is what allows the FHIR specification to retain a core level of simplicity for everyone. | [17 empty] | Element.extension | 0 | * |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  |  |  |  | n/a
DeviceAlert.derivedFrom.modifierExtension | DeviceAlert.derivedFrom.modifierExtension |  | extensions user contentmodifiers |  | 0 | * |  | Y | Y | Extension | Extensions that cannot be ignored even if unrecognized | May be used to represent additional information that is not part of the basic definition of the element and that modifies the understanding of the element in which it is contained and/or the understanding of the containing element's descendants. Usually modifier elements provide negation or qualification. To make the use of extensions safe and managable, there is a strict set of governance applied to the definition and use of extensions. Though any implementer can define an extension, there is a set of requirements that SHALL be met as part of the definition of the extension. Applications processing a resource are required to check for modifier extensions. Modifier extensions SHALL NOT change the meaning of any elements on Resource or DomainResource (including cannot change the meaning of modifierExtension itself). | There can be no stigma associated with the use of extensions by any application, project, or standard - regardless of the institution or jurisdiction that uses or defines the extensions. The use of extensions is what allows the FHIR specification to retain a core level of simplicity for everyone. | Modifier extensions allow for extensions that *cannot* be safely ignored to be clearly distinguished from the vast majority of extensions which can be safely ignored. This promotes interoperability by eliminating the need for implementers to prohibit the presence of extensions. For further information, see the [definition of modifier extensions](extensibility.html#modifierExtension). | [16 empty] | BackboneElement.modifierExtension | 0 | * |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  |  |  |  | N/A
DeviceAlert.derivedFrom.observation | DeviceAlert.derivedFrom.observation |  |  |  | 1 | 1 |  |  | Y | Reference(Observation) | The Observation having a value causing the alert condition | Indicates the Observation whose value is causing the alert condition; or, if `component` is present, the Observation with a component causing the alert condition. | [18 empty] | DeviceAlert.derivedFrom.observation | 1 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} | Event.reason
DeviceAlert.derivedFrom.component | DeviceAlert.derivedFrom.component |  |  |  | 0 | 1 |  |  |  | Coding | The Observation.component having a value causing the alert condition | If applicable, the code of the component (of the Observation identified in `derivedFrom.observation`) having a value causing the alert condition. This might be used when the alert is associated with a specific component of an Observation, rather than the overall Observation; for example, a low diastolic blood pressure. Since the component is identified by matching `Observation.component.code`, if more than one component have the same code, the specific component is ambiguous. Repetitions of this element indicate additional components contributing to the alert condition. | [10 empty] | example | Codes identifying names of simple observations. | http://hl7.org/fhir/ValueSet/observation-codes | [5 empty] | DeviceAlert.derivedFrom.component | 0 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())}
DeviceAlert.derivedFrom.limit | DeviceAlert.derivedFrom.limit |  |  |  | 0 | 1 |  |  |  | Range | The boundaries beyond which a value was detected to cause the alert condition | The limits beyond which a value was detected to cause the alert condition. The actual value is the `Observation.value[x]` referenced by `derivedFrom.observation` or, if specified, the `Observation.component.value[x]` of the component (with `Observation.component.code` matching `derivedFrom.component`) of the reference Observation. | [18 empty] | DeviceAlert.derivedFrom.limit | 0 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  |  | LimitAlertConditionState/Limits | Limit-Specification
DeviceAlert.label | DeviceAlert.label |  |  |  | 0 | 1 |  |  |  | string | Text to be displayed for the alert condition | The label might combine information from the alert code, priority, the measurement type, measurement value, body sites and other sources, e.g., “HR > 180”. | [18 empty] | DeviceAlert.label | 0 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  |  | AlertConditionDescriptor/TypeConceptDescription | AlertCondition/Label-String
DeviceAlert.signal | DeviceAlert.signal |  |  |  | 0 | * |  |  |  | BackboneElement | Annunciation or notification of the alert condition | Annunciation or notification of the alert condition. | [18 empty] | DeviceAlert.signal | 0 | * |  | ele-1:All FHIR elements must have a @value or children or both {hasValue() or (children().count() > id.count())} |  |  | AlertSignalState | AlertCondition/AlertFlags; AlertCondition/AlertControls
DeviceAlert.signal.id | DeviceAlert.signal.id |  |  |  | 0 | 1 |  |  |  | string | Unique id for inter-element referencing | Unique id for the element within a resource (for internal references). This may be any string value that does not contain spaces. | [18 empty] | Element.id | 0 | 1 | ele-1 | [5 empty] | n/a
DeviceAlert.signal.extension | DeviceAlert.signal.extension |  | extensions user content |  | 0 | * |  |  |  | Extension | Additional content defined by implementations | May be used to represent additional information that is not part of the basic definition of the element. To make the use of extensions safe and managable, there is a strict set of governance applied to the definition and use of extensions. Though any implementer can define an extension, there is a set of requirements that SHALL be met as part of the definition of the extension. | There can be no stigma associated with the use of extensions by any application, project, or standard - regardless of the institution or jurisdiction that uses or defines the extensions. The use of extensions is what allows the FHIR specification to retain a core level of simplicity for everyone. | [17 empty] | Element.extension | 0 | * |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  |  |  |  | n/a
DeviceAlert.signal.modifierExtension | DeviceAlert.signal.modifierExtension |  | extensions user contentmodifiers |  | 0 | * |  | Y | Y | Extension | Extensions that cannot be ignored even if unrecognized | May be used to represent additional information that is not part of the basic definition of the element and that modifies the understanding of the element in which it is contained and/or the understanding of the containing element's descendants. Usually modifier elements provide negation or qualification. To make the use of extensions safe and managable, there is a strict set of governance applied to the definition and use of extensions. Though any implementer can define an extension, there is a set of requirements that SHALL be met as part of the definition of the extension. Applications processing a resource are required to check for modifier extensions. Modifier extensions SHALL NOT change the meaning of any elements on Resource or DomainResource (including cannot change the meaning of modifierExtension itself). | There can be no stigma associated with the use of extensions by any application, project, or standard - regardless of the institution or jurisdiction that uses or defines the extensions. The use of extensions is what allows the FHIR specification to retain a core level of simplicity for everyone. | Modifier extensions allow for extensions that *cannot* be safely ignored to be clearly distinguished from the vast majority of extensions which can be safely ignored. This promotes interoperability by eliminating the need for implementers to prohibit the presence of extensions. For further information, see the [definition of modifier extensions](extensibility.html#modifierExtension). | [16 empty] | BackboneElement.modifierExtension | 0 | * |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  |  |  |  | N/A
DeviceAlert.signal.activationState | DeviceAlert.signal.activationState |  |  |  | 1 | 1 |  |  | Y | CodeableConcept | on ¦ off ¦ paused | Paused indicates that annunciation has temporarily been disabled ("snooze"). | [10 empty] | extensible |  | http://terminology.hl7.org/ValueSet/devicealert-activationState | [5 empty] | DeviceAlert.signal.activationState | 1 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  |  | AlertSignalState/ActivationState | AlertCondition/AlertState
DeviceAlert.signal.presence | DeviceAlert.signal.presence |  |  |  | 0 | 1 |  |  |  | CodeableConcept | on ¦ latched ¦ off ¦ ack | Indicates whether the signal is currently being annunciated. An on signal is currently being annunciated; a latched signal is currently being being annunciated although the alert condition has ended; an off signal is not currently being annunciated; and an acknowledged signal is not currently being annunciated because the user has acknowledged the signal. | [10 empty] | extensible |  | http://terminology.hl7.org/ValueSet/devicealert-presence | [5 empty] | DeviceAlert.signal.presence | 0 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  |  | AlertSignalState/Presence
DeviceAlert.signal.annunciator | DeviceAlert.signal.annunciator |  |  |  | 0 | 1 |  |  |  | CodeableReference(Device) | Where the signal is being annunciated | Signalling by the source device is local; signalling elsewhere is considered remote. A reference to the "top level" signalling device could also be present. | [10 empty] | required |  | http://terminology.hl7.org/ValueSet/devicealert-annunciation¦1.0.0 | [5 empty] | DeviceAlert.signal.annunciator | 0 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  |  | AlertSignalDescriptor/Manifestation | AlertCondition/AlertFlags
DeviceAlert.signal.manifestation | DeviceAlert.signal.manifestation |  |  |  | 0 | 1 |  |  |  | CodeableConcept | How the signal is being annunciated | How the signal is being annunciated. | [10 empty] | extensible |  | http://terminology.hl7.org/ValueSet/devicealert-manifestation | [5 empty] | DeviceAlert.signal.manifestation | 0 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  |  | AlertSignalState/Location | AlertCondition/AlertFlags
DeviceAlert.signal.type | DeviceAlert.signal.type |  |  |  | 0 | * |  |  |  | CodeableConcept | Characteristics of the signal manifestation | Details of the signal manifestation, such as a 1 meter visual indicator or a 4 meter visual indicator. | [10 empty] | example |  | http://hl7.org/fhir/ValueSet/devicealert-signalType | [5 empty] | DeviceAlert.signal.type | 0 | * |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  |  | AlertSystemDescriptor/AlertSignal/Type
DeviceAlert.signal.indication | DeviceAlert.signal.indication |  |  |  | 0 | 1 |  |  |  | Period | When the signal was being annunciated | The period during which the signal was being annunciated. If there is no indicated period end, the annunciation is on-going. | The indication period start when the signal presence is first on (or latched) and ends when the signal becomes off or acknowledged. | [17 empty] | DeviceAlert.signal.indication | 0 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  |  | Start: AlertConditionState/DeterminationTime + AlertConditionState/ActualConditionGenerationDelay + AlertSignalState/ActualSignalGenerationDelay End: Change of AlertConditionState/DeterminationTime