## Metadata
Property | Value
URL | http://hl7.org/fhir/StructureDefinition/DeviceMetric
Version | 6.0.0-ballot4
Name | DeviceMetric
Title
Status | active
Experimental | false
Date | 2026-04-01T11:29:14+00:00
Publisher | HL7 International / Devices
Contact | null (http://www.hl7.org/Special/committees/fiwg)
Contact | null (http://www.hl7.org/Special/committees/healthcaredevices)
Jurisdiction | World
Description | Describes a measurement, calculation or setting capability of a device. The DeviceMetric resource is derived from the ISO/IEEE 11073-10201 Domain Information Model standard, but is more widely applicable.
Purpose
Copyright
FHIR Version | 6.0.0-ballot4
Kind | resource
Type | DeviceMetric
Base Definition | http://hl7.org/fhir/StructureDefinition/DomainResource
Abstract | false
Derivation | specialization

## Elements
ID | Path | Slice Name | Alias(s) | Label | Min | Max | Must Support? | Is Modifier? | Is Summary? | Type(s) | Short | Definition | Comments | Requirements | Default Value | Meaning When Missing | Fixed Value | Pattern | Example | Minimum Value | Maximum Value | Maximum Length | Binding Strength | Binding Description | Binding Value Set | Code | Slicing Discriminator | Slicing Description | Slicing Ordered | Slicing Rules | Base Path | Base Min | Base Max | Condition(s) | Constraint(s) | Mapping: FiveWs Pattern Mapping | Mapping: ISO/IEEE 11073-10207: Domain Information and Service Model for Service-Oriented Point-of-Care Medical Device Communication | Mapping: ISO/IEEE 11073-10201: Domain Information Model | Mapping: ISO/IEEE 11073-10206: Personal health device communication--Abstract content information model | Mapping: RIM Mapping
DeviceMetric | DeviceMetric |  |  |  | 0 | * |  |  |  |  | Measurement, calculation or setting capability of a medical device | Describes a measurement, calculation or setting capability of a device. | The DeviceMetric resource is derived from the ISO/IEEE 11073-10201 Domain Information Model standard, but is more widely applicable. | [17 empty] | DeviceMetric | 0 | * |  | dom-2:If the resource is contained in another resource, it SHALL NOT contain nested Resources {contained.contained.empty()} dom-3:If the resource is contained in another resource, it SHALL be referred to from elsewhere in the resource or SHALL refer to the containing resource {contained.where((('#'+id.trace('id') in %resource.descendants().select(reference ¦ as(uri))) or descendants().where(reference='#' ¦ as(uri)='#').exists()).not()).trace('unmatched', id).empty()}dom-4:If a resource is contained in another resource, it SHALL NOT have a meta.versionId or a meta.lastUpdated {contained.meta.versionId.empty() and contained.meta.lastUpdated.empty()}dom-5:If a resource is contained in another resource, it SHALL NOT have a security label {contained.meta.security.empty()}dom-6:A resource should have narrative for robust management {text.`div`.exists()} | administrative.device | MetricDescriptor | Metric |  | Entity, Role, or Act
DeviceMetric.id | DeviceMetric.id |  |  |  | 0 | 1 |  |  | Y | string | Logical id of this artifact | The logical id of the resource, as used in the URL for the resource. Once assigned, this value never changes. | Within the context of the FHIR RESTful interactions, the resource has an id except for cases like the create and conditional update. Otherwise, the use of the resouce id depends on the given use case. | [17 empty] | Resource.id | 0 | 1
DeviceMetric.meta | DeviceMetric.meta |  |  |  | 0 | 1 |  |  | Y | Meta | Metadata about the resource | The metadata about the resource. This is content that is maintained by the infrastructure. Changes to the content might not always be associated with version changes to the resource. | [18 empty] | Resource.meta | 0 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())}
DeviceMetric.implicitRules | DeviceMetric.implicitRules |  |  |  | 0 | 1 |  | Y | Y | uri | A set of rules under which this content was created | A reference to a set of rules that were followed when the resource was constructed, and which must be understood when processing the content. Often, this is a reference to an implementation guide that defines the special rules along with other profiles etc. For DomainResources, if more than one implicitRules policy applies, additional policies may be conveyed using the [additionalImplicitRules](https://build.fhir.org/ig/HL7/fhir-extensions/StructureDefinition-additionalImplicitRules.html) extension. | Asserting this rule set restricts the content to be only understood by a limited set of trading partners. This inherently limits the usefulness of the data in the long term. However, the existing health eco-system is highly fractured, and not yet ready to define, collect, and exchange data in a generally computable sense. Wherever possible, implementers and/or specification writers should avoid using this element. Often, when used, the URL is a reference to an implementation guide that defines these special rules as part of its narrative along with other profiles, value sets, etc. | [17 empty] | Resource.implicitRules | 0 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())}
DeviceMetric.language | DeviceMetric.language |  |  |  | 0 | 1 |  |  |  | code | Language of the resource content | The base language in which the resource is written. | Language is provided to support indexing and accessibility (typically, services such as text to speech use the language tag). The html language tag in the narrative applies to the narrative. The language tag on the resource may be used to specify the language of other presentations generated from the data in the resource. Not all the content has to be in the base language. The Resource.language should not be assumed to apply to the narrative automatically. If a language is specified, it should it also be specified on the div element in the html (see rules in HTML5 for information about the relationship between xml:lang and the html lang attribute). | [9 empty] | required | IETF language tag for a human language | http://hl7.org/fhir/ValueSet/all-languages¦6.0.0-ballot4 | [5 empty] | Resource.language | 0 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())}
DeviceMetric.text | DeviceMetric.text |  | narrative htmlxhtmldisplay |  | 0 | 1 |  |  |  | Narrative | Text summary of the resource, for human interpretation | A human-readable narrative that contains a summary of the resource and can be used to represent the content of the resource to a human. The narrative need not encode all the structured data, but is required to contain sufficient detail to make it "clinically safe" for a human to just read the narrative. Resource definitions may define what content should be represented in the narrative to ensure clinical safety. | Contained resources do not have a narrative. Resources that are not contained SHOULD have a narrative. In some cases, a resource may only have text with little or no additional discrete data (as long as all minOccurs=1 elements are satisfied). This may be necessary for data from legacy systems where information is captured as a "text blob" or where text is additionally entered raw or narrated and encoded information is added later. | [17 empty] | DomainResource.text | 0 | 1 | dom-6 | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  |  |  |  | Act.text?
DeviceMetric.contained | DeviceMetric.contained |  | inline resources anonymous resourcescontained resources |  | 0 | * |  |  |  | Resource | Contained, inline Resources | These resources do not have an independent existence apart from the resource that contains them - they cannot be identified independently, nor can they have their own independent transaction scope. This is allowed to be a Parameters resource if and only if it is referenced by a resource that provides context/meaning. | This should never be done when the content can be identified properly, as once identification is lost, it is extremely difficult (and context dependent) to restore it again. Contained resources may have profiles and tags in their meta elements, but SHALL NOT have security labels. Contained resources may be a resource type defined in the FHIR specification, or an [additional resource](resource.html#additional). | [17 empty] | DomainResource.contained | 0 | * | dom-2 dom-4dom-3dom-5 | [5 empty] | N/A
DeviceMetric.extension | DeviceMetric.extension |  | extensions user content |  | 0 | * |  |  |  | Extension | Additional content defined by implementations | May be used to represent additional information that is not part of the basic definition of the resource. To make the use of extensions safe and managable, there is a strict set of governance applied to the definition and use of extensions. Though any implementer can define an extension, there is a set of requirements that SHALL be met as part of the definition of the extension. | There can be no stigma associated with the use of extensions by any application, project, or standard - regardless of the institution or jurisdiction that uses or defines the extensions. The use of extensions is what allows the FHIR specification to retain a core level of simplicity for everyone. | [17 empty] | DomainResource.extension | 0 | * |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  |  |  |  | N/A
DeviceMetric.modifierExtension | DeviceMetric.modifierExtension |  | extensions user content |  | 0 | * |  | Y | Y | Extension | Extensions that cannot be ignored | May be used to represent additional information that is not part of the basic definition of the resource and that modifies the understanding of the element that contains it and/or the understanding of the containing element's descendants. Usually modifier elements provide negation or qualification. To make the use of extensions safe and managable, there is a strict set of governance applied to the definition and use of extensions. Though any implementer is allowed to define an extension, there is a set of requirements that SHALL be met as part of the definition of the extension. Applications processing a resource are required to check for modifier extensions. Modifier extensions SHALL NOT change the meaning of any elements on Resource or DomainResource (including cannot change the meaning of modifierExtension itself). | There can be no stigma associated with the use of extensions by any application, project, or standard - regardless of the institution or jurisdiction that uses or defines the extensions. The use of extensions is what allows the FHIR specification to retain a core level of simplicity for everyone. | Modifier extensions allow for extensions that *cannot* be safely ignored to be clearly distinguished from the vast majority of extensions which can be safely ignored. This promotes interoperability by eliminating the need for implementers to prohibit the presence of extensions. For further information, see the [definition of modifier extensions](extensibility.html#modifierExtension). | [16 empty] | DomainResource.modifierExtension | 0 | * |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  |  |  |  | N/A
DeviceMetric.identifier | DeviceMetric.identifier |  |  |  | 0 | * |  |  | Y | Identifier | Instance identifier | Instance identifiers assigned to a device, by the device or gateway software, manufacturers, other organizations or owners. For example, handle ID. | For identifiers assigned to a device by the device or gateway software, the `system` element of the identifier should be set to any identifier of the device. | [17 empty] | DeviceMetric.identifier | 0 | * |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} | FiveWs.identifier
DeviceMetric.status | DeviceMetric.status |  |  |  | 1 | 1 |  | Y | Y | code | active ¦ inactive ¦ entered-in-error ¦ unknown | Indicates current state of the record. | Unknown does not represent "other" - one of the defined statuses must apply. Unknown is used when the authoring system is not sure what the current status is. | [9 empty] | required | Describes the status of the DeviceMetric. | http://hl7.org/fhir/ValueSet/metric-status¦6.0.0-ballot4 | [5 empty] | DeviceMetric.status | 1 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} | FiveWs.status
DeviceMetric.operationalStatus | DeviceMetric.operationalStatus |  |  |  | 0 | 1 |  |  | Y | code | on ¦ off ¦ standby ¦ unknown | Indicates current operational state of the device. For example: On, Off, Standby, etc. |  |  |  | If absent, the operational status of the device is unknown. | [6 empty] | required | Describes the operational status of the DeviceMetric. | http://hl7.org/fhir/ValueSet/metric-operational-status¦6.0.0-ballot4 | [5 empty] | DeviceMetric.operationalStatus | 0 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  | Metric/ActivationState | Metric::Metric-Status
DeviceMetric.category | DeviceMetric.category |  |  |  | 1 | 1 |  |  | Y | CodeableConcept | The kind of metric represented | Indicates the category of the observation generation process. A DeviceMetric can be, for example, a setting, measurement, or calculation. | [10 empty] | extensible | Describes the category of the metric. | http://hl7.org/fhir/ValueSet/metric-category | [5 empty] | DeviceMetric.category | 1 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} | FiveWs.class | MetricDescriptor/MetricCategory | Metric::Metric-Specification::category | Filled in when this can be derived from SystemInfo.observation-type-list[i]
DeviceMetric.type | DeviceMetric.type |  |  |  | 1 | 1 |  |  | Y | CodeableConcept | Identity of metric, for example Heart Rate or PEEP Setting | Describes the type of the metric. For example: Heart Rate, PEEP Setting, etc. | `DeviceMetric.type` can be referred to either IEEE 11073-10101 or LOINC. | [9 empty] | preferred | IEEE 11073-10101 | http://hl7.org/fhir/ValueSet/devicemetric-type | [5 empty] | DeviceMetric.type | 1 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} | FiveWs.what[x] | MetricDescriptor/Type | type = VMO::Type type.text = VMO::Label-String | SystemInfo.observation-type-list[i]
DeviceMetric.device | DeviceMetric.device |  |  |  | 1 | 1 |  |  | Y | Reference(Device) | The device to which this DeviceMetric applies | The device to which this DeviceMetric applies. The referenced device could be a top-level device or a component of a device. | If the reference is not to a top-level device, the top-level device is reached through the device's `parent` element (possibly recursively). | [17 empty] | DeviceMetric.device | 1 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} | FiveWs.subject
DeviceMetric.unit | DeviceMetric.unit |  |  |  | 0 | 1 |  |  | Y | CodeableConcept | Unit of Measure for the Metric | Describes the unit that an observed value determined for this metric will have. For example: Percent, Seconds, etc. | `DeviceMetric.unit` can refer to either [UCUM](https://terminology.hl7.org/UCUM.html) or preferably an [RTMMS](https://terminology.hl7.org/MDC.html) coding system. | [9 empty] | preferred | IEEE 11073-10101 | http://hl7.org/fhir/ValueSet/ucum-units | [5 empty] | DeviceMetric.unit | 0 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  | MetricDescriptor/Unit | unit = Metric::Unit-Code unit.text = Unit::Label-String | UCUM unit used for SystemInfo.observation-type-list[i] observations
DeviceMetric.color | DeviceMetric.color |  |  |  | 0 | 1 |  |  |  | code | Color name (from CSS4) or #RRGGBB code | The preferred color associated with the metric (e.g., display color). This is often used to aid clinicians to track and identify parameter types by color. In practice, consider a Patient Monitor that has ECG/HR and Pleth; the metrics are displayed in different characteristic colors, such as HR in blue, BP in green, and PR and SpO2 in magenta. |  |  |  | If absent, there is no preferred color for rendering. | [6 empty] | required | Describes the typical color of representation. | http://hl7.org/fhir/ValueSet/color-codes¦6.0.0-ballot4 | [5 empty] | DeviceMetric.color | 0 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  |  | Metric::Color
DeviceMetric.measurementFrequency | DeviceMetric.measurementFrequency |  |  |  | 0 | 1 |  |  |  | Quantity | Indicates how often the metric is taken or recorded | The frequency at which the metric is taken or recorded. Devices measure metrics at a wide range of frequencies; for example, an ECG might sample measurements in the millisecond range, while an NIBP might trigger only once an hour. Less often, the measurementFrequency is based on a unit other than time, such as distance (e.g., for a measuring wheel). The update period could be different than the measurement frequency, if the device does not update the published observed value with the same frequency as it was measured. | [18 empty] | DeviceMetric.measurementFrequency | 0 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  | MetricDescriptor/DeterminationPeriod | value = Metric::Measure-Period::value or Metric::Metric-Specification::update-period unit = Metric::Measure-Period::unit-code
DeviceMetric.availability | DeviceMetric.availability |  |  |  | 0 | 1 |  |  | Y | CodeableConcept | The continuity of the metric (e.g., measurement) | Indicates whether the metric is derived (measured) by a continuous means or by an intermittent means. | For example, a gauge on a pressurized vessel showing current pressure is a continuous measurement, while a thermometer that produces readings on demand is an intermittent measurement. Note that metric availability is typically a static characteristic and is distinct from the operational status. | [9 empty] | extensible | Describes the availability of the metric | http://hl7.org/fhir/ValueSet/metric-availability | [5 empty] | DeviceMetric.availability | 0 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  | MetricDescriptor/MetricAvailability
DeviceMetric.calibration | DeviceMetric.calibration |  |  |  | 0 | * |  |  |  | BackboneElement | Describes the calibrations that have been performed or that are required to be performed | Describes the calibrations that have been performed or that are required to be performed. | [18 empty] | DeviceMetric.calibration | 0 | * |  | ele-1:All FHIR elements must have a @value or children or both {hasValue() or (children().count() > id.count())}
DeviceMetric.calibration.id | DeviceMetric.calibration.id |  |  |  | 0 | 1 |  |  |  | string | Unique id for inter-element referencing | Unique id for the element within a resource (for internal references). This may be any string value that does not contain spaces. | [18 empty] | Element.id | 0 | 1 | ele-1 | [5 empty] | n/a
DeviceMetric.calibration.extension | DeviceMetric.calibration.extension |  | extensions user content |  | 0 | * |  |  |  | Extension | Additional content defined by implementations | May be used to represent additional information that is not part of the basic definition of the element. To make the use of extensions safe and managable, there is a strict set of governance applied to the definition and use of extensions. Though any implementer can define an extension, there is a set of requirements that SHALL be met as part of the definition of the extension. | There can be no stigma associated with the use of extensions by any application, project, or standard - regardless of the institution or jurisdiction that uses or defines the extensions. The use of extensions is what allows the FHIR specification to retain a core level of simplicity for everyone. | [17 empty] | Element.extension | 0 | * |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  |  |  |  | n/a
DeviceMetric.calibration.modifierExtension | DeviceMetric.calibration.modifierExtension |  | extensions user contentmodifiers |  | 0 | * |  | Y | Y | Extension | Extensions that cannot be ignored even if unrecognized | May be used to represent additional information that is not part of the basic definition of the element and that modifies the understanding of the element in which it is contained and/or the understanding of the containing element's descendants. Usually modifier elements provide negation or qualification. To make the use of extensions safe and managable, there is a strict set of governance applied to the definition and use of extensions. Though any implementer can define an extension, there is a set of requirements that SHALL be met as part of the definition of the extension. Applications processing a resource are required to check for modifier extensions. Modifier extensions SHALL NOT change the meaning of any elements on Resource or DomainResource (including cannot change the meaning of modifierExtension itself). | There can be no stigma associated with the use of extensions by any application, project, or standard - regardless of the institution or jurisdiction that uses or defines the extensions. The use of extensions is what allows the FHIR specification to retain a core level of simplicity for everyone. | Modifier extensions allow for extensions that *cannot* be safely ignored to be clearly distinguished from the vast majority of extensions which can be safely ignored. This promotes interoperability by eliminating the need for implementers to prohibit the presence of extensions. For further information, see the [definition of modifier extensions](extensibility.html#modifierExtension). | [16 empty] | BackboneElement.modifierExtension | 0 | * |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  |  |  |  | N/A
DeviceMetric.calibration.type | DeviceMetric.calibration.type |  |  |  | 0 | 1 |  |  |  | CodeableConcept | The method of calibration | Describes the type of the calibration method. | [10 empty] | extensible | Describes the type of a metric calibration. | http://hl7.org/fhir/ValueSet/metric-calibration-type | [5 empty] | DeviceMetric.calibration.type | 0 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  |  | Metric::Metric-Calibration::cal-type
DeviceMetric.calibration.state | DeviceMetric.calibration.state |  |  |  | 0 | 1 |  |  |  | code | not-calibrated ¦ calibration-required ¦ calibrated ¦ unspecified | Describes the state of the calibration. |  |  |  | If absent, the calibration state is unknown. | [6 empty] | required | Describes the state of a metric calibration. | http://hl7.org/fhir/ValueSet/metric-calibration-state¦6.0.0-ballot4 | [5 empty] | DeviceMetric.calibration.state | 0 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  |  | Metric::Metric-Calibration::cal-state
DeviceMetric.calibration.time | DeviceMetric.calibration.time |  |  |  | 0 | 1 |  |  |  | instant | Describes the time last calibration has been performed | Describes the time last calibration has been performed. | [18 empty] | DeviceMetric.calibration.time | 0 | 1 |  | ele-1:All FHIR elements must have a @value or children {hasValue() or (children().count() > id.count())} |  |  | Metric::Metric-Calibration::cal-time